Instructions for calibrating the Zebra force sensor

mounted on the CRS A460 arm at U.B.C. LCI

John Lloyd, Aug  8, 1996

Tools required:

3/32 allen driver

Force sensor calibration is a bit tedious but hopefully can at least be

made easier by these instructions.

1. Power up and calibrate the CRS arm, and clear all objects from the

workspace, and activate the RCCL environment.

2. Move the arm to the position 0,-80,170,0,0,0 using the RCCL

program 'move':

% move to 0,-80,170,0,0,0

This position, or something similar, will make it easier to attach the

force calibration jig.

3. Detach the gripper by unplugging it's cord from link 4 of the robot,

and loosening the three allen screws from the round flange at the top

of the gripper. The gripper plus the round flange should now come

away from the force sensor assembly.

4. Remove the gripper from the round flange by loosening the 4 allen

screws on the top face of the flange.

5. Using the same 4 allen screws, mount the calibration jig onto the

flange, on the same side which was occupied by the gripper, so that

the +x axis indicated on the jig lies between the +x and +y axes

(indicated by red markings) on the flange.

6. Now reattach the flange to the force sensor assembly, using the

original three allen screws, so that the +y axis indicated on the

flange is aligned with the +y axis indicated on the sensor assembly.

*  Each arm of the jig contains 4 threaded holes, spaced 5 cm apart.

On each arm, one of these holes is occupied by a piece of threaded

stock which acts as a precise stop for placing weights on the arm.

Stops can therefore be placed 5, 10, 15, or 20 cm from the center

of the jig.

7. Place all stops into the holes 15 cm from the center of the jig.

8. Start the RCCL program forceCal.

*  The calibration procedure uses 6 different weights:

Name	Mass(kg)   1/2 thickness(mm)

rsmall     1.085   6.875

rmed     2.263   8.5

rbig     3.428   9.375

gsmall     1.092   6.875

gmed     2.319   8.4

gbig     3.420   9.375

Weights with names starting with 'g' are color coded green, while those

with names starting with 'r' are color coded red.

9. Execute the script forceCalDefs.ctm, located in $RCCL/src/programs:

FORCECAL> @forceCalDefs.ctm

This will enter information about the weights into forceCal.

10. Enter the command

FORCECAL> move to position forcecal1

After the arm has come to rest, enter the command

FORCECAL> reset bias

11. Enter the commands

FORCECAL> set d1Offset 152.3

FORCECAL> set d2Offset -152.3

These reflect the distances imposed by the stop settings (with the extra

2.3 mm accounting for the stop thickness).

*  Weights will now be placed onto the jig arms. When doing so, the

indicated weight should be slid over the indicated arm until it

comes to rest at the stop.

12. Start data logging with command

FORCECAL> open output-file <fileName>

13. Place gsmall onto the +x arm and rsmall onto the -x arm, and enter

FORCECAL> enter data x-axis

Mass1, d1, mass2, d2 (units: kg,mm)> gsmall rsmall

Place gmed onto the +x arm and rmed onto the -x arm, and enter

FORCECAL> enter data x-axis

Mass1, d1, mass2, d2 (units: kg,mm)> gmed rmed

Place gbig onto the +x arm and rbig onto the -x arm, and enter

FORCECAL> enter data x-axis

Mass1, d1, mass2, d2 (units: kg,mm)> gbig rbig

Place rsmall onto the -x arm, and enter

FORCECAL> enter data x-axis

Mass1, d1, mass2, d2 (units: kg,mm)> 0 0 rsmall

Place gsmall onto the +x arm, and enter

FORCECAL> enter data x-axis

Mass1, d1, mass2, d2 (units: kg,mm)> gsmall 0 0

14. Enter the command

FORCECAL> move to position forcecal2

After the arm has come to rest, enter the command

FORCECAL> reset bias

15. Place rsmall onto the -y arm, and enter

FORCECAL> enter data y-axis

Mass1, d1, mass2, d2 (units: kg,mm)> 0 0 rsmall

Place gsmall onto the +y arm, and enter

FORCECAL> enter data y-axis

Mass1, d1, mass2, d2 (units: kg,mm)> gsmall 0 0

16. Enter the command

FORCECAL> move to position forcecal3

After the arm has come to rest, enter the command

FORCECAL> reset bias

17. Move all the stops so that they are 5 cm from the center of the

jig. Then enter the comand

FORCECAL> set d1Offset 52.3

FORCECAL> set d2Offset -52.3

18. Place gsmall onto the +x arm and rsmall onto the -x arm, and enter

FORCECAL> enter data x-axis

Mass1, d1, mass2, d2 (units: kg,mm)> gsmall rsmall

Place gmed onto the +x arm and rmed onto the -x arm, and enter

FORCECAL> enter data x-axis

Mass1, d1, mass2, d2 (units: kg,mm)> gmed rmed

Place gbig onto the +x arm and rbig onto the -x arm, and enter

FORCECAL> enter data x-axis

Mass1, d1, mass2, d2 (units: kg,mm)> gbig rbig

Place gbig onto the +x arm, and enter

FORCECAL> enter data x-axis

Mass1, d1, mass2, d2 (units: kg,mm)> gbig 0 0

Place rbig onto the -x arm, and enter

FORCECAL> enter data x-axis

Mass1, d1, mass2, d2 (units: kg,mm)> 0 0 rbig

19. Enter the command

FORCECAL> move to position forcecal4

After the arm has come to rest, enter the command

FORCECAL> reset bias

20. Place gsmall onto the +y arm and rsmall onto the -y arm, and enter

FORCECAL> enter data y-axis

Mass1, d1, mass2, d2 (units: kg,mm)> gsmall rsmall

Place gmed onto the +y arm and rmed onto the -y arm, and enter

FORCECAL> enter data y-axis

Mass1, d1, mass2, d2 (units: kg,mm)> gmed rmed

Place gbig onto the +y arm and rbig onto the -y arm, and enter

FORCECAL> enter data y-axis

Mass1, d1, mass2, d2 (units: kg,mm)> gbig rbig

21. Exit the program with the command

FORCECAL> quit

The log file <fileName> should now be input into 'matlab',

where the command f/d will produce the calibration matrix.

The contents of this calibration matrix should be entered

as the values for

static float matrix[6][6]

in $RCCL/src/puma/crs_comfxns.c. Then do 'make libs' in

$RCCL/src/puma, followed by 'make bins' in $RCCL/src. Any

application programs which use the force sensor will also have

to be recompiled.